## Hoja1
Oscilloscope mode
Vertical
Frequency response - dc coupled | Without probes and test leads (with BB120) | 123B: dc to 20 MHz (-3 dB)
124B and 125B: dc to 40 MHz (-3 dB)
| With STL120-IV 1:1 shielded test leads | DC to 12.5 MHz (-3 dB) / dc to 20 MHz (-6 dB)
| With VP41 10:1 Probe | 123B: dc to 20MHz (-3 dB)
124B and 125B: dc to 40 MHz (-3 dB)
Frequency response - ac coupled (lf roll off) | Without probes and test leads | <10 Hz (-3 dB)
| With STL120-IV 1:1 shielded test leads | <10 Hz (-3 dB)
| With VP41 10:1 Probe | <10 Hz (-3 dB)
Rise time, excluding probes, | 123B <17.5 ns
test leads | 124B and 125B <8.75 ns
Input impedance | Without probes and test leads | 1 MΩ//20 pF
| With BB120 | 1 MΩ//24 pF
| With STL120-IV 1:1 shielded test leads | 1 MΩ//230 pF
| With VP41 10:1 Probe | 5 MΩ//15.5 pF
Sensitivity | 5 mV to 200 V/div
Analog bandwidth limiter | 10 kHz
Display modes | A, -A, B, -B
Max. input voltage A and B | Direct, with test leads, or with VP41 Probe | 600 Vrms Cat IV, 750 Vrms maximum voltage.
| With BB120 | 600 Vrms
Max. floating voltage, from any terminal to ground | 600 Vrms Cat IV, 750 Vrms up to 400Hz
Horizontal
Scope modes | Normal, Single, Roll
Ranges (normal) | Equivalent sampling | 123B: 20 ns to 500 ns/div,
124B and 125B: 10 ns to 500 ns/div
| Real time sampling | 1 μs to 5 s/div
| Single (real time) | 1 μs to 5 s/div
| Roll (real time) | 1s to 60 s/div
Sampling rate (for both channels simultaneously) | Equivalent sampling (repetitive signals) | Up to 4 GS/s
| Real time sampling 1 μs to 60 s/div | 40 MS/s
Trigger
Screen update | Free run, on trigger
Source | A, B
Sensitivity A and B | @ DC to 5 MHz | 0.5 divisions or 5 mV
| @ 40 MHz | 123B: 4 divisions
124B and 125B: 1.5 divisions
| @ 60 MHz | 123B: N/A
124B and 125B: 4 divisions
Slope | Positive, negative
Advanced scope functions
Display modes | Normal | Captures up to 25 ns glitches and displays analog-like persistence waveform
| Smooth | Suppresses noise from a waveform
| Glitch off | Does not capture glitches between samples
| Envelope | Records and displays the minimum and maximum of waveforms over time
Auto set (Connect-and-View™) | Continuous fully automatic adjustments of amplitude, time base, trigger levels, trigger gap, and hold-off. Manual override by user adjustment of amplitude, time base, or trigger level.
Dual input meter
The accuracy of all measurements is within ±(% of reading + number of counts) from 18 °C to 28 °C.
Add 0.1x (specific accuracy) for each °C below 18 °C or above 28 °C. For voltage measurements with 10:1 probe, add probe uncertainty +1%. More than one waveform period must be visible on the screen.
Input A and input B
DC voltage (VDC)
Ranges | 500 mV, 5 V, 50 V, 500 V, 750 V
Accuracy | ±(0.5% +5 counts)
Common mode rejection (CMRR) | >100 dB @ dc, >60 dB @ 50, 60, or 400 Hz
Full scale reading | 5000 counts
True-rms voltages (V ac and V ac+dc)
Ranges | 500 mV, 5 V, 50 V, 500 V, 750 V
Accuracy for 5% to 100% of range (DC coupled) | DC to 60 Hz (V ac+dc) | ±(1% +10 counts)
| 1 Hz to 60 Hz (V ac) | ±(1% +10 counts)
Accuracy for 5% to 100% of range (AC or dc coupled) | 60 Hz to 20 kHz | ±(2.5% +15 counts)
DC rejection (only VAC) | >50 dB
Common mode rejection (CMRR) | >100 dB @ dc
>60 dB @ 50, 60, or 400 Hz
Full scale reading | 5000 counts, reading is independent of any signal crest factor.
Peak
Modes | Max peak, Min peak, or pk-to-pk
Ranges | 500 mV, 5 V, 50 V, 500 V, 2200 V
Accuracy | Accuracy Max peak or Min peak | 5% of full scale
| Accuracy Peak-to-Peak | 10% of full scale
Full scale reading | 500 counts
Frequency (Hz)
Ranges | 123B: 1 Hz, 10 Hz, 100 Hz, 1 kHz, 10 kHz, 100 kHz,1 MHz, 10 MHz, and 50 MHz
124B and 125B: 1 Hz, 10 Hz, 100 Hz, 1 kHz, 10 kHz, 100 kHz, 1 MHz, 10 MHz, and 70 MHz
Frequency range | 15 Hz (1 Hz) to 50 MHz in continuous autoset
Accuracy @1 Hz to 1 MHz | ±(0.5% +2 counts)
Full scale reading | 10,000 counts
RPM
Max reading | 50.00 kRPM
Accuracy | ±(0.5% +2 counts)
Duty cycle (PULSE)
Range | 2% to 98%
Frequency range | 15 Hz (1 Hz) to 30 MHz in continuous autoset
Pulse width (PULSE)
Frequency range | 15 Hz (1 Hz) to 30 MHz in continuous autoset
Full scale reading | 1000 counts
Amperes (AMP)
With current clamp | Ranges | Same as V dc, V ac, V ac+dc, or PEAK
| Scale factors | 0.1 mV/A, 1 mV/A, 10 mV/A, 100 mV/A, 400 mV/A, 1 V/A, 10 mV/mA
| Accuracy | Same as V dc, V ac, V ac+dc, or PEAK (add current clamp uncertainty)
Temperature (TEMP) with optional temperature probe
Range | 200 °C/div (200 °F/div)
Scale factor | 1 mV/°C and 1 mV/°F
Accuracy | As V dc (add temp. probe uncertainty)
Decibel (dB)
0 dBV | 1 V
0 dBm (600 Ω /50 Ω) | 1 mW referenced to 600 Ω or 50 Ω
dB on | V dc, V ac, or Vac+dc
Full scale reading | 1000 counts
Crest factor (CREST)
Range | 1 to 10
Full scale reading | 90 Counts
Phase
Modes | A to B, B to A
Range | 0 to 359 degrees
Resolution | 1 degree
Power (125B only)
Configurations | 1 phase / 3 phase 3 conductor balanced loads (3 phase: fundamental component only, AUTOSET mode only)
Power factor (PF) | Ratio between watts and VA range - 0.00 to 1.00
Watt | RMS reading of multiplying corresponding samples of input A (volts) and input B (amperes)
| Full scale reading | 999 counts
VA | Vrms x Arms
| Full scale reading | 999 counts
VA reactive (var) | √((VA)2-W2)
| Full scale reading | 999 counts
Vpwm
Purpose | To measure on pulse width modulated signals, like motor drive inverter outputs
Principle | Readings show the effective voltage based on the average value of samples over a whole number of periods of the fundamental frequency
Accuracy | As Vrms for sinewave signals
Input A to common
Ohm (Ω)
Ranges | 123B and 124B | 500 Ω , 5 kΩ, 50 kΩ, 500 kΩ, 5 MΩ, 30 MΩ
| 125B | 50 Ω, 500 Ω , 5 kΩ, 50 kΩ, 500 kΩ, 5 MΩ, 30 MΩ
Accuracy | ±(0.6% + 5 counts) 50 Ω ±(2% + 20 counts)
Full scale reading | 50 Ω to 5 MΩ - 5000 counts, 30 MΩ - 3000 counts
Measurement current | 0.5 mA to 50 nA, decreases with increasing ranges
Open circuit voltage | <4 V
Continuity (Cont)
Beep | <(30 Ω ±5 Ω) in 50 Ω range
Measurement current | 0.5 mA
Detection of shorts of | ≥1 ms
Diode
Measurement voltage | @0.5 mA | >2.8 V
| @open circuit | <4 V
Measurement current | 0.5 mA
Polarity | + on input A, - on COM
Capacitance (CAP)
Ranges | 50 nF, 500 nF, 5 μF, 50 μF, 500 μF
Full scale reading | 5000 counts
Measurement current | 500 nA to 0.5 mA, increases with increasing ranges
Advanced meter functions
Zero Set | Set actual value to reference
AutoHold (on A) | Captures and freezes a stable measurement result. Beeps when stable. AutoHold works on the main meter reading, with thresholds of 1 Vpp for AC signals and 100 mV for DC signals.
Fixed decimal point | Activated by using attenuation keys
Cursor Readout (124B and 125B)
Sources | A, B
Single vertical line | Average, min and max readout
Average, min, max and time from start of readout (in ROLL mode; instrument in HOLD)
Min, max and time from start of readout (in RECORDER mode; instrument in HOLD)
Harmonics values in POWER QUALITY mode.
Dual vertical lines | Peak-peak, time distance and reciprocal time distance readout
Average, min, max and time distance readout (in ROLL mode; instrument in HOLD)
Dual horizontal lines | High, low and peak-peak readout
Rise or fall time | Transition time, 0%-level and 100%-level readout (manual or auto leveling; auto leveling only
possible in single channel mode)
Accuracy | As oscilloscope accuracy
Recorder
The recorder captures meter readings in Meter Recorder mode or continuously captures waveform samples in Scope Recorder mode. The information is stored on internal memory or on optional SD card (with the 125B or 124B).
The results are displayed as Chart recorder display that plots a graph of min and max values of Meter measurements over time or as a waveform recorder display that plots all the captured samples.
Meter readings
Measurement Speed | Maximum 2 measurement/s
Record Size (min, max, average) | 2 M readings for 1 channel
Recorded Time Span | 2 weeks
Maximum number of events | 1024
Waveform record
Maximum sample rate | 400 K sample/s
Size Internal memory | 400 M samples Recorded Time
Span internal memory | 15 minutes at 500 μs/div
11 hours at 20 ms/div
Record Size SD card | 1.5 G samples
Recorded Time Span SD card | 11 hours at 500 μs/div
14 days at 20 ms/div
Maximum number of events | 64
Power Quality (125B only)
Readings | Watt, VA, var, PF, DPF, Hz
Watt, VA, var ranges (auto) | 250 W to 250 MW, 625 MW, 1.56 GW
| When selected: total (% r) | ±(2% + 6 counts)
| When selected: fundamental (% f) | ±(4% + 4 counts)
DPF | 0.00 to 1.00
PF | 0.00 to 1.00, ±0.04
Frequency range | 10.0 Hz to 15.0 kHz
40.0 Hz to 70.0 Hz
Number of Harmonics | DC to 51
Readings / Cursor readings | Vrms / Arms /Watt | Each harmonic from fundamental maybe selected for individual readings
(fundamental 40 Hz to 70 Hz)
Includes frequency of fundamental, phase Angle and K-factor (in Amp and Watt)
BusHealth tester (Fluke 125B only)
Type | Subtype | Protocol
AS-i | NEN-EN50295
CAN | ISO-11898
Interbus S | RS-422 | EIA-422
Modbus | RS-232 | RS-232/EIA-232
| RS-485 | RS-485/EIA-485
Foundation Fieldbus | H1 | 61158 type 1, 31.25 kBit
Profibus | DP | EIA-485
| PA | 61158 type 1
Miscellaneous
Display | Type | 5.7-inch color active matrix TFT
| Resolution | 640 x 480 pixels
Waveform Display | Vertical | 10 div of 40 pixels
| Horizontal | 12 div of 40 pixels
Power | External | Via Power Adapter BC430
| Input voltage | 10 V DC to 21 V DC
| Power consumption | 5 W typical
| Input connector | 5 mm jack
| Internal | Via Battery Pack BP290
| Battery power | Rechargeable Li-Ion 10.8 V
| Operating time | 7 hours with 50% backlight brightness
| Charging time | 4 hours with test tool off, 7 hours with test tool on
| Allowable ambient temp | 0 to 40 °C (32 to 104 °F) during charging
Memory | Internal memory can store 20 data sets (screen waveform and setup) | Micro SD card slot with optional SD card (max size of 32 GB)
Mechanical | Size | 259 x 132 x 55 mm
(10.2 x 5.2 x 2.15 in)
| Weight | 1.4 kg (3.2 lb) including battery pack
Interface | Optically isolated | Transfer screen copies (bitmaps), settings and data
| USB to PC/laptop | OC4USB optically isolated USB adapter/cable, (optional), using FlukeView® software for Windows®.
| Optional WiFi adapter | Fast transfer of screen copies (bitmaps), settings and data to PC/laptop, tablet, smartphone, etc. A USB port is provided for attaching the WiFi dongle. Do not use the USB port with a cable for safety reasons.
Environmental
Environmental | MIL-PRF-28800F, Class 2
Temperature | Battery Operation | 0 to 40 °C (32 to 104 °F)
| Power Adapter Operation | 0 to 50 °C (32 to 122 °F)
| Storage | -20 to 60 °C (-4 to 140 °F)
Humidity (Operating) | @ 0 to 10 °C (32 to 50 °F) | Non-condensing
| @ 10 to 30 °C (50 to 86 °F) | 95%
| @ 30 to 40 °C (86 to 104 °F) | 75%
| @ 40 to 50 °C (104 to 122 °F) | 45%
Storage | @ -20 to 60 °C (-4 to 140 °F) | Non-condensing
Altitude | Operating at 3 km (10,000 feet) | CAT III 600 V
| Operating at 2 km (6,600 feet) | CAT IV 600 V
| Storage | 12 km (40,000 feet)
EMC electromagnetic compatibility | International | IEC 61326-1: Industrial, CISPR 11: Group 1, Class A
| Korea (KCC) | Class A Equipment (Industrial Broadcasting & Communication Equipment)
| USA (FCC) | 47 CFR 15 subpart B. This product is considered an exempt device per clause 15.103.
Wireless radio with adapter | Frequency range | 2412 MHz to 2462 MHz
| Output power | <100 mW
Enclosure protection | IP51, ref: EN/IEC60529
Safety | General | IEC 61010-1: Pollution Degree 2
| Measurement | IEC 61010-2-033: CAT IV 600 V/CAT III 750 V
Max. input voltage input A and B | Direct on input or with leads | 600 Vrms CAT IV for derating
| With Banana-to-BNC Adapter BB120 | 600 Vrms for derating
| Max. floating voltage from any terminal to ground | 600 Vrms Cat IV, 750 Vrms up to 400 Hz