## Hoja1
Specifications
Accuracy
Parameter |  | Range | Resolution | Intrinsic Accuracy at Reference Conditions (% of Reading +% of Full Scale)
Voltage |  | 1000 V | 0.1 V | ±(0.2% + 0.01%)
Current: Direct input | i17xx-flex 1500 12" | 150 A | 0.1 A | ±(1% + 0.02%)
|  | 1500 A | 1:00 | ±(1% + 0.02%)
| i17xx-flex 3000 24" | 300 A | 1:00 | ±(1% + 0.03%)
|  | 3000 A | 10:00 | ±(1% + 0.03%)
| i17xx-flex 6000 36" | 600 A | 1:00 | ±(1.5% + 0.03%)
|  | 6000 A | 10:00 | ±(1.5% + 0.03%)
| i40s-EL clamp | 4:00 | 1 mA | ±(0.7% + 0.02%)
|  | 40 A | 10 mA | ±(0.7% + 0.02%)
Frequency |  | 42.5 Hz to 69 Hz | 0.01 Hz | ±(0.1%)
Auxillary input |  | ±10 V dc | 0.1 mV | ±(0.2% + 0.02%)
Voltage min/max |  | 1000 V | 0.1 V | ±(1% + 0.1%)
Current min/max |  | Defined by accessory | Defined by accessory | ±(5% + 0.2%)
THD on voltage |  | 1000% | 0.1% | ±0.5
THD on current |  | 1000% | 0.1% | ±0.5
Voltage harmonics 2nd to 50th |  | 1000 V | 0.1 V | ≥ 10 V: ±5% of reading
< 10 V: ±0.5V
Current harmonics 2nd to 50th |  | Defined by accessory | Defined by accessory | ≥ 3% of current range: ±5% of reading
< 3% of current range: ±0.15% of range
Unbalance |  | 100% | 0.1% | ±0.2
Intrinsic Uncertainty ±(% of reading +% of range)¹
Parameter | Influence quantity | iFlex1500-12 | iFlex3000-24 | iFlex6000-36 | i40S-EL
|  | 150 A / 1500 A | 300 A / 3000 A | 600 A / 6000 A | 4 A / 40 A
Active Power P | PF ≥ 0.99 | 1.2% + 0.005% | 1.2% + 0.0075% | 1.7% + 0.0075% | 1.2% + 0.005%
Active Energy Ea | PF ≥ 0.99 | 1.2% + 0.005% | 1.2% + 0.0075% | 1.7% + 0.0075% | 1.2% + 0.005%
Apparent power S | 0 ≤ PF ≤ 1 | 1.2% + 0.005% | 1.2% + 0.0075% | 1.7% + 0.0075% | 1.2% + 0.005%
Apparent Energy Eap | 0 ≤ PF ≤ 1 | 1.2% + 0.005% | 1.2% + 0.0075% | 1.7% + 0.0075% | 1.2% + 0.005%
Reactive power Q | 0 ≤ PF ≤ 1 | 2.5% of measured apparent power
Reactive Energy Er | 0 ≤ PF ≤ 1 | 2.5% of measured apparent power
Power Factor PF | - | ±0.025
Displacement Power Factor | - | ±0.025
DBF/cosϕ | - | ±0.025
Additional uncertainty in% of range¹ | VP-N>250 V | 0.015% | 0.0225% | 0.0225% | 0.015%
¹Range = 1000 V x Irange
Reference conditions:
Environmental: 23 °C ±5 °C, instrument operating for at least 30 minutes, no external electrical/magnetic field, RH <65%
Input conditions: Cosϕ/PF=1, Sinusoidal signal f=50 Hz/60 Hz, power supply 120 V/230 V ±10%.
Current and power specifications: Input voltage 1 ph: 120 V/230 V or 3 ph wye/delta: 230 V/400 V Input current: I > 10% of Irange
Primary conductor of clamps or Rogowski coil in center position
Temperature coefficient: Add 0.1 x specified accuracy for each degree C above 28 °C or below 18 °C
Electrical Specifications
Power Supply
Voltage range | 100 V to 500 V using safety plug input when powering from the measurement circuit
100 V to 240 V using standard power cord (IEC 60320 C7)
Power consumption | Maximum 50 VA (max. 15 VA when powered using IEC 60320 input)
Efficiency | ≥ 68.2% (in accordance with energy efficiency regulations)
Maximum no-load consumption | < 0.3 W only when powered using IEC 60320 input
Mains power frequency | 50/60 Hz ±15%
Battery | Li-ion 3.7 V, 9.25 Wh, customer-replaceable
On-battery runtime | Four hours in standard operating mode, up to 5.5 hours in power saving mode
Charging time | < 6 hours
Data Acquisition
Resolution | 16-bit synchronous sampling
Sampling frequency | 10.24 kHz at 50/60 Hz, synchronized to mains frequency
Input signal frequency | 50/60 Hz (42.5 to 69 Hz)
Circuit types | 1-φ, 1-φ IT, Split phase, 3-φ delta, 3-φ wye, 3-φ wye IT, 3-φ wye balanced, 3-φ Aron/Blondel (2-element delta), 3-φ delta open leg, Currents only (load studies)
Data storage | Internal flash memory (not user replaceable)
Memory size | Typical 10 logging sessions of 8 weeks with 1-minute intervals and 500 events¹
¹The number of possible logging sessions and logging period depends on user requirements.
Basic Interval
Measured parameters | Voltage, current, aux, frequency, THD V, THD A, power, power factor, fundamental power, DPF, energy
Averaging interval | User selectable: 1 sec, 5 sec, 10 sec, 30 sec, 1 min, 5 min, 10 min, 15 min, 30 min
Averaging time min/max values | Voltage, Current: Full cycle RMS updated every half cycle (URMS1/2 according to IEC61000-4-30 Aux, Power: 200ms
Demand Interval (Energy Meter Mode)
Measured parameters | Energy (Wh, varh, VAh), PF, maximum demand, cost of energy
Interval | User selectable: 5 min, 10 min, 15 min, 20 min, 30 min, off
Power Quality Measurements
Measured parameter | Voltage, frequency, unbalance, voltage harmonics, THD V, current, harmonics, THD A, TDD
Averaging interval | 10 min
Individual harmonics | 2nd to 50th harmonic
Total harmonic distortion | Calculated on 50 harmonics
Events | Voltage: dips, swells, interruptions, current: inrush current
Triggered recordings | Full cycle RMS updated every half cycle of voltage and current (Urms1/2 according to IEC61000-4-30)
Waveform of voltage and current (1738)
Standards Compliance
Harmonics | IEC 61000-4-7: Class 1
IEEE 519 (short time harmonics)
Power quality | IEC 61000-4-30 Class S, IEC62586-1 (PQI-S device)
Power | IEEE 1459
Power quality compliance | EN50160 (for measured parameters)
Interfaces
USB-A | File transfer via USB flash drive, firmware updates
Max. current: 120 mA
WiFi | File transfer and remote control via direct connection or WiFi infrastructure
Bluetooth | Read auxiliary measurement data from Fluke Connect® 3000 series modules (requires 1738 or 1736 upgrade option)
USB-mini | Data download device to PC
Voltage Inputs
Number of inputs | 4 (3 phases and neutral)
Maximum input voltage | 1000 Vrms, CF 1.7
Input impedance | 10 MΩ
Bandwidth (-3 dB) | 42.5 Hz - 3.5 kHz
Scaling | 1:1 and variable
Measurement category | 1000 V CAT III/600 V CAT IV
Current Inputs
Number of inputs | 4, mode selected automatically for attached sensor
Input voltage | Clamp input: 500 mVrms/50 mVrms; CF 2.8
Rogowski coil input | 150 mVrms/15 mVrms at 50 Hz, 180 mVrms/18 mVrms at 60 Hz; CF 4; all at nominal probe range
Range | 1 A to 150 A/10 A to 1500 A with thin flexible current probe i17XX-flex1500 12"
3 A to 300 A/30 A to 3000 A with thin flexible current probe i17XX-flex3000 24"
6 A to 600 A/60 A to 6000 A with thin flexible current probe i17XX-flex6000 36"
40 mA to 4 A/0.4 A to 40 A with 40 A clamp i40s-EL
Bandwidth (-3 dB) | 42.5 Hz - 3.5 kHz
Scaling | 1:1 and variable
Auxiliary Inputs
Number of inputs | 2
Input range | 0 to ±10 V dc, 1reading/s
Scale factor | Format: mx + b (gain and offset) user configurable
Displayed units | User configurable (7 characters, for example, °C, psi, or m/s)
Wireless Connection
Number of inputs | 2
Supported modules | Fluke Connect® 3000 series
Acquisition | 1 reading/s
Environmental Specifications
Operating temperature | -10 °C to +50 °C (14 °F to 122 °F)
Storage temperature | -20 °C to +60 °C (-4 °F to 140 °F), with battery: -20 °C to +50 °C (-4 °F to 122 °F)
Operating humidity | 10 °C to 30 °C (50 °F to 86 °F) max. 95% RH
30 °C to 40 °C (86 °F to 104 °F) max. 75% RH
40 °C to 50 °C (104 °F to 122 °F) max. 45% RH
Operating altitude | 2000 m (up to 4000 m derate to 1000 V CAT II/600 V CAT III/300 V CAT IV)
Storage altitude | 12,000m
Enclosure | IP50 in accordance with EN60529
Vibration | MIL-T-28800E, Type 3, Class III, Style B
Safety | IEC 61010-1
IEC Mains Input: Overvoltage Category II, Pollution Degree 2
Voltage Terminals: Overvoltage Category IV, Pollution Degree 2
IEC 61010-2-031: CAT IV 600 V / CAT III 1000 V
Electromagnetic compatibility | EN 61326-1: Industrial CISPR 11: Group 1, Class A
Korea (KCC): Class A Equipment (industrial broadcasting and communication equipment)
USA (FCC): 47 CFR 15 subpart B. This product is considered an exempt device per clause 15.103
Temperature coefficient | 0.1 x accuracy specification/°C
General Specifications
Color LCD display | 4.3-inch active matrix TFT, 480 pixels x 272 pixels, resistive touch panel
Warranty | Instrument and power supply: Two-years (battery not included)
Accessories: one-year
Calibration cycle: two-years
Dimensions | Instrument: 19.8 cm x 16.7 cm x 5.5 cm (7.8 in x 6.6 in x 2.2 in)
Power supply: 13.0 cm x 13.0 cm x 4.5 cm (5.1 in x 5.1 in x 1.8 in)
Instrument with power supply attached: 19.8 cm x 16.7 cm x 9 cm (7.8 in x 6.6 in x 3.5 in)
Weight | Instrument: 1.1 kg (2.5 lb)
Power supply: 400 g (0.9 lb)
Tamper protection | Kensington lock slot
i17xx-flex 1500 12" Flexible Current Probe Specifications
Measuring range | 1 to 150 A ac/10 to 1500 A ac
Nondestructive current | 100 kA (50/60 Hz)
Intrinsic error at reference condition* | ±0.7% of reading
Accuracy 173x + iFlex | ±(1% of reading + 0.02% of range)
Temperature coefficient over operating temperature range | 0.05% of reading/°C 0.09% of reading/°F
Working voltage | 1000 V CAT III, 600 V CAT IV
Probe cable length | 305 mm (12 in)
Probe cable diameter | 7.5 mm (0.3 in)
Minimum bending radius | 38 mm (1.5 in)
Output cable length | 2 m (6.6 ft)
Weight | 115 g
Probe cable material | TPR
Coupling material | POM + ABS/PC
Output cable | TPR/PVC
Operating temperature | -20 °C to +70 °C (-4 °F to 158 °F) temperature of conductor under test shall not exceed 80 °C (176 °F)
Temperature, non-operating | -40 °C to +80 °C (-40 °F to 176 °F)
Relative humidity, operating | 15% to 85% non-condensing
IP rating | IEC 60529:IP50
Warranty | One-year
* Reference condition:
Environmental: 23 °C ±5 °C, no external electrical/magnetic field, RH 65%
Primary conductor in center position