## Hoja1
General Specifications
AC Current via Jaw
Range | 600.0 A
Resolution | 0.1 A
Accuracy | 2% ±5 digits (10 Hz to 100 Hz)
2.5% ±5 digits (100-500 Hz)
Crest Factor (50 Hz/60 Hz) | 2.5 @ 600 A
Add 2% for C.F. >2
AC Current via Flexible Current Probe
Range | 2500 A
Resolution | 0.1 A (≤ 600 A)
1 A (≤ 2500 A)
Accuracy | 3% ±5 digits (5 – 500 Hz)
Crest Factor (50/60Hz) | 2.5 at 1400 A
1.42 at 2500 A
Add 2% for C.F. > 2
Position Sensitivity
Distance from Optimum
|  | i2500-10 Flex | i2500-18 Flex | Error
| A | 0.5 in (12.7 mm) | 1.4 in (35.6 mm) | ±0.5 %
| B | 0.8 in (20.3 mm) | 2.0 in (50.8 mm) | ±1.0 %
| C | 1.4 in (35.6 mm) | 2.5 in (63.5 mm) | ±2.0 %
Measurement uncertainty assumes centralized primary conductor at optimum position, no external electrical or magnetic field, and within operating temperature range.
DC Current
Range | 600.0 A
Resolution | 0.1 A
Accuracy | 2% ±5 digits
AC Voltage
Range | 1000 V
Resolution | 0.1 V (≤600.0 V)
1 V (≤1000 V)
Accuracy | 1.5% ±5 digits (20 Hz to 500 Hz)
DC Voltage
Range | 1000 V
Resolution | 0.1 V (≤600.0 V)
1 V (≤1000 V)
Accuracy | 1% ±5 digits
mV dc
Range | 500.0 mV
Resolution | 0.1 mV
Accuracy | 1% ±5 digits
Resistance
Range | 6000 Ω
Resolution | 0.1 Ω (≤600 Ω)
1 Ω (≤6000 Ω)
Accuracy | 1% ±5 digits
Capacitance
Range | 1000 μF
Resolution | 0.1 μF (≤ 100 μF)
1 μ F (≤ 1000 μF)
Accuracy | 1% ±4 digits
Mechanical Specifications
Maximum voltage between any terminal and earth ground | 1000 V
Batteries | 2 AA, NEDA 15A, IEC LR6
Operating temperature | -10 °C to +50 °C
Storage temperature | -40 °C to +60 °C
Operating humidity - | ≤90% RH (at 10 °C to 30 °C)
non-condensing (< 10°C) | ≤75% RH (at 30 °C to 40 °C)
≤45% RH (at 40 °C to 50 °C)
Operating altitude | 3000 m
Storage altitude | 12,000 m
Size (L x W x H) | 249 x 85 x 45 mm
Weight | 395 g
Jaw opening | 34 mm
Flexible current probe diameter | 7.5 mm
Flexible current probe cable length | 1.8 m
(head to electronics connector)
Safety | IEC 61010-1, Pollution Degree 2
IEC 61010-2-032: CAT III 1000 V / CAT IV 600 V
IEC 61010-2-033:CAT III 1000 V / CAT IV 600 V
IP rating | IEC 60529: IP30, non-operating
Radio Frequency Certification FCC ID | T68-FBLE IC:6627A-FBLE
Electromagnetic Compatibility (EMC)
International | IEC 61326-1: Portable, Electromagnetic Environment, IEC 61326-2-2
CISPR 11: Group 1, Class A
Group 1: Equipment has intentionally generated and/or uses conductively-coupled radio frequency energy that is necessary for the internal function of the equipment itself.
Class A: Equipment is suitable for use in all establishments other than domestic and those directly connected to a low-voltage power supply network that supplies buildings used for domestic purpose. There may be potential difficulties in ensuring electromagnetic compatibility in other environments due to conducted and radiated disturbances. Emissions that exceed the levels required by CISPR 11 can occur when the equipment is connected to a test object.
Korea (KCC) | Class A Equipment (Industrial Broadcasting & Communication Equipment)
Class A: Equipment meets requirements for industrial electromagnetic wave equipment and the seller or user should take notice of it. This equipment is intended for use in business environments and not to be used in homes.
USA (FCC) | 47 CFR 15 subpart B. This product is considered an exempt device per clause 15.103.
Temperature coefficients | Add 0.1 x specified accuracy for each degree C above 28 °C or below 18 °C