## Hoja1
Electrical Specifications
AC Current via Jaw
Range | 999.9 A
Resolution | 0.1 A
Accuracy | 2% ± 5 digits (10-100 Hz)
5% ± 5 digits (100-500 Hz)
Crest Factor (50/60 Hz) | 3 @ 500 A
2.5 @ 600 A
1.42 @ 1000 A
Add 2% for C.F. > 2
AC Current via Flexible Current Probe
Range | 999.9 A / 2500 A (45 Hz – 500 Hz)
Resolution | 0.1 A / 1 A
Accuracy | 3% ± 5 digits
Crest Factor (50/60 Hz) | 3.0 at 1100 A
2.5 at 1400 A
1.42 at 2500 A
Add 2% for C.F. > 2
Position Sensitivity
Distance from Optimum
|  | i2500-10 Flex | i2500-18 Flex | Error
| A | 0.5 in (12.7 mm) | 1.4 in (35.6 mm) | ±0.5%
| B | 0.8 in (20.3 mm) | 2.0 in (50.8 mm) | ±1.0%
| C | 1.4 in (35.6 mm) | 2.5 in (63.5 mm) | ±2.0%
Measurement uncertainty assumes centralized primary conductor at optimum position, no external electrical or magnetic field, and within operating temperature range.
DC Current
Range | 999.9 A
Resolution | 0.1 A
Accuracy | 2% ± 5 digits
AC Voltage
Range | 600 V /1000 V
Resolution | 0.1 V / 1 V
Accuracy | 1.5% ± 5 digits (20 – 500 Hz)
DC Voltage
Range | 600.0 V /1000 V
Resolution | 0.1 V / 1 V
Accuracy | 1% ± 5 digits
Frequency – Via Jaw
Range | 5.0 – 500.0 Hz
Resolution | 0.1 Hz
Accuracy | 0.5% ± 5 digits
Trigger Level | 5 – 10 Hz, ≥ 10 A
10 – 100 Hz, ≥ 5 A
100 – 500 Hz, ≥ 10 A
Frequency via Flexible Current Probe
Range | 5.0 to 500.0 Hz
Resolution | 0.1 Hz
Accuracy | 0.5% ± 5 digits
Trigger Level | 5 to 20 Hz, ≥ 25 A
20 to 100 Hz, ≥ 20 A
100 to 500 Hz, ≥ 25 A
Resistance
Range | 600 Ω/6 kΩ/60 kΩ
Resolution | 0.1 Ω/1 Ω/10 Ω
Accuracy | 1% ± 5 digits
Mechanical Specifications
Size (L x W x H) | 277 x 88 x 43 mm (55 mm for remote unit)
Weight | 350 g
Jaw Opening | 34 mm
Flexible Current Probe Diameter | 7.5 mm
Flexible Current Probe Cable Length | 1.8 m
(head to electronics connector)
Environmental Specifications
Operating Temperature | -10°C to +50°C
Storage Temperature | -40°C to +60°C
Operating Humidity | Non condensing (< 10°C)
≤ 90% RH (at 10°C to 30°C)
≤ 75% RH (at 30°C to 40°C)
≤ 45% RH (at 40°C to 50°C)
(Without Condensation)
Operating Altitude | 2,000 meters
Storage Altitude | 12,000 meters
EMI, RFI, EMC, RF | EN 61326-1:2006, EN 61326-2-2:2006
ETSI EN 300 328 V1.7.1:2006
ETSI EN 300 489 V1.8.1:2008
FCC Part 15 Subpart C Sections 15.207, 15.209, 15.249
FCCID: T68-F381
RSS-210 IC: 6627A-F381
Temperature Coefficients | Add 0.1 x specified accuracy for each degree C above 28°C or below 18°C
Wireless Frequency | 2.4 GHz ISM Band 10 meter range
Safety Compliance | ANSI/ISA S82.02.01:2004
CAN/CSA-C22.2 No. 61010-1-04
IEC/EN 61010-1:2001 to 1000V CAT III, 600V CAT IV.
Double Insulation Clearance | Per IEC 61010-2-032
Double Insulation Creepage | Per IEC 61010-1
Agency Approvals