## Hoja1
Measurement Accuracy
Voltage DC | 30.000 V | 0.01% + 2 mV (upper display)
| 20.000 V | 0.01% + 2 mV (lower display)
| 90.00 mV | 0.01% + 10 μV
| -10.00 mV to 75.00 mV | 0.025% + 1 count (via TC connector)
Current DC | 24.000 mA | 0.01%+ 2 mA
Resistance | 0.0 to 4000.0 Ω | ± 0.015%
Frequency | 2.0 to 1000.0 CPM | 0.05% + 0.1 cpm
| 1.0 to 1000.0 Hz | 0.05% + 0.1 Hz
| 1.00 to 15.00 kHz | 0.05% + 0.01 kHz
| Sensitivity | 1 V peak-to-peak-minimum
Pressure | Accuracy from 0.025% of range using any of 29 pressure modules (for detailed specifications refer to pressure modules in options and accessories). Modules available for differential, gauge, vacuum, absolute, dual and high pressure.
Source Accuracy
Voltage DC | 100.00 mV | 0.01% + 10 μV
| 20.000 V | 0.01% + 2 mV
| -10.00 mV to 75.00 mV | 0.025% + 1 count (via TC connector)
Current DC | 24.000 mA (Source) | 0.01% + 2 μA
| 24.000 mA (Simulate) | 0.01% + 2 μA
Resistance | Range, accuracy (percent of reading + floor) and excitation current | 5 to 400 Ω, 0.015% + 0.1 Ω, 0.1 to 0.5 mA
5 to 400 Ω, 0.015% + 0.05 Ω, 0.5 to 3 mA
400 to 1500 Ω, 0.015% + 0.5 Ω, 0.05 to 0.8 mA
1500 to 4000 Ω, 0.015% + 0.5 Ω, 0.05 to 0.4 mA
Frequency | 2.0 to 1000.0 CPM | 0.05%
| 1.0 to 1000.0 Hz | 0.05%
| 1.0 to 10.0 kHz | 0.25%
| 10.0 to 15.0 kHz | 0.5%
| Waveform | 5 V p-p squarewave, -0.1 V offset
RTDs and Thermocouples
Measure and source accuracy | Cu10 | 1.8°C
| NI-120 | 0.15°C
| PT-100 (385) | 0.45°C
| PT-100 (392) | 0.35°C
| PT-100 (391-JIS) | 0.35°C
| PT-200 (385) | 0.95°C
| PT-500 (385) | 0.55°C
| PT-1000 (385) | 0.35°C
| RTD Resolution | 0.01°C (0.01°F)
| Cu10.01 C | 0.1°C (0.1°F)
| J | 0.4°C
| K | 0.5°C
| T | 0.4°C
| E | 0.4°C
| R | 1.4°C
| S | 1.4°C
| B | 1.4°C
| L | 0.45°C
| U | 0.45°C
| N | 0.6°C
| XK | 0.4°C
| BP | 1.1°C
| TC Resolution | J, K, T, E, L, N, U: 0.1°C, 0.1°F B, R, S: 1°C, 1°F
Specifications
Ramp functions | Source functions | Voltage, current, resistance, frequency, temperature
| Ramps | Slow ramp, Fast ramp, 25% step-ramp
Loop power function | Voltage | 24 V
| Accuracy | 10%
| Maximum current | 22 mA, short circuit protected
Step functions | Source functions | Voltage, current, resistance, frequency, temperature
| Steps | 25% of range, 100% of range
Environmental Specifications
Operating temperature | -10°C to +55°C
Storage temperature | -20°C to +71°C
Operating altitude | 3000 m
Safety Specifications
Agency approvals | EN 61010-1:1993, ANSI/ISA S82.01-1994; CAN/CSA C22.2 No 1010.1:1992
Mechanical and General Specifications
Size | 200 x 96 x 47 mm (7.9 x 3.8 x 1.9 in)
Weight | 650 g (23 oz)
Batteries | 4 AA alkaline batteries
Battery replacement | Separate battery compartment, accessible without breaking calibration seal
Side port connections | Pressure module connector, also used for remote real-time programming