## Hoja1
Measurement Accuracy
Voltage DC | Range/ Resolution | 1 Year |  | 2 Years
| 100.000 mV | 0.02% + 0.005 mV |  | 0.03% + 0.005 mV
| 3.00000 V | 0.02% + 0.00005 V |  | 0.03% + 0.00005 V
| 30.0000 V | 0.02% + 0.0005 V |  | 0.03% + 0.0005 V
| 300.00 V | 0.05% + 0.05 V |  | 0.07% + 0.05 V
Voltage AC | 3.000 V (40 Hz to 500 Hz) / 0.001 V | 0.5% + 0.002 V |  | 1.0% + 0.004 V
| 30.00 V (40 Hz to 500 Hz) / 0.01 V | 0.5% + 0.02 V |  | 1.0% + 0.04 V
| 300.0 V (40 Hz to 500 Hz) / 0.1 V | 0.5% + 0.2 V |  | 1.0% + 0.2 V
Current DC | 30.000 mA | 0.01% + 5 uA |  | 0.015% + 7 uA
| 110.00 mA | 0.01% + 20 uA |  | 0.015% + 30 uA
Resistance | 10.000 Ω | 0.05% + 50 mΩ |  | 0.07% + 70 mΩ
| 100.00 Ω | 0.05% + 50 mΩ |  | 0.07% + 70 mΩ
| 1.0000 kΩ | 0.05% + 500 mΩ |  | 0.07% + 0.5 Ω
| 10.000 kΩ | 0.1% + 10 Ω |  | 0.15% + 15 Ω
Frequency | 1.00 to 110.00 Hz / 0.01 Hz |  |  | 0.05 Hz
| 110.1 to 1100.0 Hz / 0.1 Hz |  |  | 0.5 Hz
| 1.101 to 11.000 kHz / 0.001 kHz |  |  | 0.005 kHz
| 11.01 to 50.00 kHz / 0.01 kHz |  |  | 0.05 kHz
Source Accuracy
|  | 1 Year |  | 2 Years
Voltage DC | 100.000 mV | 0.01% + 0.005 mV |  | 0.015% + 0.005 mV
| 1.00000 V | 0.01% + 0.00005 V |  | 0.015% + 0.0005 V
| 15.0000 V | 0.01% + 0.0005 V |  | 0.015% + 0.0005 V
Current DC | 22.000 mA (source) | 0.01% + 0.003 mA |  | 0.02% + 0.003 mA
| Current sink (simulate) | 0.02% + 0.007 mA |  | 0.04% + 0.007 mA
Resistance | 10.000 Ω | 0.01% + 10 mΩ |  | 0.015% + 15 mΩ
| 100.00 Ω | 0.01% + 20 mΩ |  | 0.015% + 30 mΩ
| 1.0000 kΩ | 0.02% + 0.2 Ω |  | 0.03% + 0.3 Ω
| 10.000 kΩ | 0.02% + 3 Ω |  | 0.03% + 5 Ω
Frequency | 0.1 to 10.99 Hz |  |  | 0.01 Hz
| 0.01 to 10.99 Hz |  |  | 0.01 Hz
| 11.00 to 109.99 Hz |  |  | 0.1 Hz
| 110.0 to 1099.9 Hz |  |  | 0.1 Hz
| 1.100 to 21.999 kHz |  |  | 0.002 kHz
| 22.000 to 50.000 kHz |  |  | 0.005 kHz
Technical Data
Data log functions | Measure functions | Voltage, current, resistance, frequency, temperature, pressure
| Reading rate | 1, 2, 5, 10, 20, 30, or 60 readings/minute
| Maximum record length | 8000 readings (7980 for 30 or 60 readings/minute)
| Ramp functions | Source functions | Voltage, current, resistance, frequency, temperature
|  | Rate | 4 steps/second
|  | Trip detect | Continuity or voltage (continuity detection not available when sourcing current)
Loop power function | Voltage | Selectable, 26 V
| Accuracy | 10%, 18 V minimum at 22 mA
| Maximum current | 25 mA, short circuit protected
| Maximum input voltage | 50 V DC
Step functions | Source functions | Voltage, current, resistance, frequency, temperature
| Manual step | Selectable step, change with arrow buttons
| Autostep | Fully programmable for function, start delay, stepvalue, time per step, repeat
Environmental Specifications
Operating temperature | -10°C to +50°C
Storage temperature | -20°C to +60°C
Dust/water resistance | Meets IP52, IEC 529
Operating altitude | 3000 m above mean sea level (9842 ft)
Safety Specifications
Agency approvals | CAN/CSA C22.2 No 1010.1-92, ASNI/ISA S82.01-1994, UL3111, and EN610-1:1993
Mechanical and General Specifications
Size | 136 x 245 x 63 mm (5.4 x 9.6 x 2.5 in)
Weight | 1.2 kg (2.7 lb)
Batteries | Internal Battery Pack Li-ion: 7.2V,4400mAh, 30 Wh
Battery life | >8 hours typical
Battery replacement | Replace without opening calibrator; no tools required
Side port connections | Pressure module connector
USB Connector to interface to your PC
Digital instrument (HART) connector
Connection for optional battery charger/eliminator
Data storage capacity | 1 week of calibration procedures results
90 day specifications | The standard specification interval for the 750 Seriesare 1 and 2 years.
Typical 90 day measurement and source accuracy can be estimated by dividing the one year "% of reading" or "% of output" specifications by 2.
Floor specifications, expressed as "% of full scale" or "counts" or "ohms" remain constant.
Temperature, Resistance Temperature Detectors
Degrees or % of reading - Type (α) | Range °C | Measure °C1
|  | 1 year |  | 2 year
100 Ω Pt (385) | -200 to 100 | 0.07°C |  | 0.14°C
| 100 to 800 | 0.02% + 0.05°C |  | 0.04% + 0.10°C
200 Ω Pt (385) | -200 to 100 | 0.07°C |  | 0.14°C
| 100 to 630 | 0.02% + 0.05°C |  | 0.04% + 0.10°C
500 Ω Pt (385) | -200 to 100 | 0.07°C |  | 0.14°C
| 100 to 630 | 0.02% + 0.05°C |  | 0.04% + 0.10°C
1000 Ω Pt (385) | -200 to 100 | 0.07°C |  | 0.14°C
| 100 to 630 | 0.02% + 0.05°C |  | 0.04% + 0.10°C
100 Ω Pt (3916) | -200 to 100 | 0.07°C |  | 0.14°C
| 100 to 630 | 0.02% + 0.05°C |  | 0.04% + 0.10°C
100 Ω Pt (3926) | -200 to 100 | 0.08°C |  | 0.16°C
| 100 to 630 | 0.02% + 0.06°C |  | 0.04% + 0.12°C
10 Ω Cu (427) | -100 to 260 | 0.2°C |  | 0.4°C
120 Ω Ni (672) | -80 to 260 | 0.1°C |  | 0.2°C
Source current | Source °C |  |  | Allowable current2
| 1 year | 2 year
1 mA | 0.05°C | 0.10°C |  | 0.1 mA to 10 mA
| 0.0125% + 0.04°C | 0.025% + 0.08°C
500 μA | 0.06°C | 0.12°C |  | 0.1 mA to 1 mA
| 0.017% + 0.05°C | 0.034% + 0.10°C
250 μA | 0.06°C | 0.12°C |  | 0.1 mA to 1 mA
| 0.017% + 0.05°C | 0.034% + 0.10°C
150 μA | 0.06 C | 0.12 C |  | 0.1 mA to 1 mA
| 0.017% + 0.05°C | 0.034% + 0.10°C
1 mA | 0.05°C | 0.10°C |  | 0.1 mA to 10 mA
| 0.0125% + 0.04°C | 0.025% + 0.08°C
1 mA | 0.05°C | 0.10°C |  | 0.1 mA to 10 mA
| 0.0125% + 0.04°C | 0.025% + 0.08°C
3 mA | 0.2°C | 0.4°C |  | 0.1 mA to 10 mA
1 mA | 0.04°C | 0.08°C |  | 0.1 mA to 10 mA
1. For two and three-wire RTD measurements, add 0.4°C to the specifications.
2. Supports pulsed transmitters and PLCs with pulse times as short as 1 ms
Temperature, Thermocouples
Type | Source °C | Measure °C |  | Source °C
|  | 1 year | 2 years | 1 year | 2 years
E | -250 to -200 | 1.3 | 2.0 | 0.6 | 0.9
| -200 to -100 | 0.5 | 0.8 | 0.3 | 0.4
| -100 to 600 | 0.3 | 0.4 | 0.3 | 0.4
| 600 to 1000 | 0.4 | 0.6 | 0.2 | 0.3
N | -200 to -100 | 1.0 | 1.5 | 0.6 | 0.9
| -100 to 900 | 0.5 | 0.8 | 0.5 | 0.8
| 900 to 1300 | 0.6 | 0.9 | 0.3 | 0.4
J | -210 to -100 | 0.6 | 0.9 | 0.3 | 0.4
| -100 to 800 | 0.3 | 0.4 | 0.2 | 0.3
| 800 to 1200 | 0.5 | 0.8 | 0.3 | 0.3
K | -200 to -100 | 0.7 | 1.0 | 0.4 | 0.6
| -100 to 400 | 0.3 | 0.4 | 0.3 | 0.4
| 400 to 1200 | 0.5 | 0.8 | 0.3 | 0.4
| 1200 to 1372 | 0.7 | 1.0 | 0.3 | 0.4
T | -250 to -200 | 1.7 | 2.5 | 0.9 | 1.4
| -200 to 0 | 0.6 | 0.9 | 0.4 | 0.6
| 0 to 400 | 0.3 | 0.4 | 0.3 | 0.4
B | 600 to 800 | 1.3 | 2.0 | 1.0 | 1.5
| 800 to 1000 | 1.0 | 1.5 | 0.8 | 1.2
| 1000 to 1820 | 0.9 | 1.3 | 0.8 | 1.2
R | -20 to 0 | 2.3 | 2.8 | 1.2 | 1.8
| 0 to 100 | 1.5 | 2.2 | 1.1 | 1.7
| 100 to 1767 | 1.0 | 1.5 | 0.9 | 1.4
S | -20 to 0 | 2.3 | 2.8 | 1.2 | 1.8
| 0 to 200 | 1.5 | 2.1 | 1.1 | 1.7
| 200 to 1400 | 0.9 | 1.4 | 0.9 | 1.4
| 1400 to 1767 | 1.1 | 1.7 | 1.0 | 1.5
C | 0 to 800 | 0.6 | 0.9 | 0.6 | 0.9
| 800 to 1200 | 0.8 | 1.2 | 0.7 | 1.0
| 1200 to 1800 | 1.1 | 1.6 | 0.9 | 1.4
| 1800 to 2316 | 2.0 | 3.0 | 1.3 | 2.0
L | -200 to -100 | 0.6 | 0.9 | 0.3 | 0.4
| -100 to 800 | 0.3 | 0.4 | 0.2 | 0.3
| 800 to 900 | 0.5 | 0.8 | 0.2 | 0.3
U | -200 to 0 | 0.6 | 0.9 | 0.4 | 0.6
| 0 to 600 | 0.3 | 0.4 | 0.3 | 0.4
BP | 0 to 1000 | 1.0 | 1.5 | 0.4 | 0.6
| 1000 to 2000 | 1.6 | 2.4 | 0.6 | 0.9
| 2000 to 2500 | 2.0 | 3.0 | 0.8 | 1.2
XK | -200 to 300 | 0.2 | 0.3 | 0.2 | 0.5
| 300 to 800 | 0.4 | 0.6 | 0.3 | 0.6